## Slide 1: PIGE in SIMNRA 7
M. Mayer
Max-Planck-Institut für Plasmaphysik, Garching, Germany

## Slide 2: Introduction: Simulation of IBA spectra
Monte-Carlo simulation: SRIM, CORTEO, MCERD, GEANT, …
Trajectories of individual particles
High level of detail: Multiple scattering, geometry, …
Long computing times
Computing time improvements introduce additional problems
Analytical codes: SIMNRA, RUMP, NDF, …
Energy distributions of ensembles of particles
Trajectories approximated as straight lines + corrections: energy spread by multiple scattering, geometry, …
Very fast computation
In many practical cases comparable level of detail as MC
In some cases approximations get invalid

## Slide 3: RBS/NRA/ERDA and PIGE in SIMNRA
RBS/NRA/ERDA
1. Start at surface with initial ensemble - Window can be included
2. Move ensemble into the sample - Propagate energy distribution (mean energy, straggling)
3. Scattering event or reaction: - Number of outgoing particles from cross-section - Initial energy distribution from kinematics
4. Move ensemble to surface Propagate energy distribution (mean energy, straggling)
5. Propagation through foil + detector
6. Sort into channels

## Slide 4: RBS/NRA/ERDA and PIGE in SIMNRA
PIGE
1. Start at surface with initial ensemble - Window can be included
2. Move ensemble into the sample - Propagate energy distribution (mean energy, straggling)
3. Scattering event or reaction: - Number of outgoing particles from cross-section
4. No absorption of outgoing particles

## Slide 5: Input data SIMNRA 7.04
Atomic masses: Recommended values of the 2012 update to the atomic mass evaluation Wang 2012
Isotopic abundances: Recommended values of isotopic abundances De Bievre 1993
Physical constants: CODATA recommended values 2018 Tiesinga 2021
Data interpolation: Linear interpolation for stopping powers and cross-sections  Least informative approach without additional knowledge  SRIM stopping powers are not smooth due to internal linear interpolation

## Slide 6: Calculation of Energy Loss
SIMNRA since 5.70
5th order Runge-Kutta method
Embedded 4th order method with Cash-Karp parameters for error estimate and step-width control
No derivatives of stopping power required
Accuracy per single Runge-Kutta step < 1 eV
 Accuracy ~10-5 for ‘Auto’ step width control  Accuracy ~10-6 for step width of 10 keV or less
 Accuracy eV to few 10 eV for energy losses of MeV
Error relative to 1 keV step width
SIMNRA 7.04

## Slide 7: Calculation of Energy Spread
1st moment: Mean energy  Stopping power: ZBL, SRIM,…
2nd moment: Straggling  Straggling model: Bohr, Chu, Yang
3rd moment: Skewness  Free Coulomb scattering
SIMNRA 7:
3rd moment taken into account for all energy spread components
Energy distributions approximated by two-piece normal distributions (TPND’s)
Free Coulomb scattering for 3rd moment
Tschalär effect always included (includes 3rd moment)
19.68 MeV protons in Al
Tschalär and Maccabee, Phys. Rev. B 1 (1970) 2863

## Slide 8: Energy Spread of Protons in Si
Approximation as Gaussian only valid for energy loss > 10 – 20%
Sharp resonances close to surface require taking asymmetry into account
See exercise #2
2 MeV protons in Si
Initial beam spread 1 keV
Approximation by SIMNRA
hup, hlow: HWHM to low (high) energies
Asymmetry

## Slide 9: Multiple small-angle scattering
Small energy losses
Almost symmetric distribution with non-Gaussian tails
Large energy losses
Asymmetric distribution due to path length differences
High-energy edge approaches Gaussian
Long tail towards low-energies due to nuclear straggling
2 MeV 4He in Au, a = 60°
De
a
d

## Slide 10: Multiple small-angle scattering and number of emitted photons
Multiple scattering changes the path length in the material
Mean path length = Path length weighted with multiple scattering distribution
In 3 dimensions
Example: Exercise 5 (11.41 µm Si / 0.98 µm Al / Si)
Multiple scattering in Al layer - At entrance 3.6° FWHM - At exit 3.8° FWHM
Analytical result: 36.88732 Counts
SIMNRA, no Multiple Scattering: 36.88657 -20.3 ppm
SIMNRA, with Multiple Scattering: 36.89688 +259 ppm
d

## Slide 11: Calculation of the Number of Reaction Products
For resonant cross-sections straggling distribution has to be taken into account
Number of created Gammas in a specific depth = Integral of (Straggling*Cross-section)
Integral calculated in [E0 – 4s, E0 + 4s] + correction ~0.14%
5 point Gauss-Mehler integration between 2 data points, max integration range s
Accuracy 10-7 for constant cross-sections
12C(p,p)12C
Straggling
Straggling*Cross-section

## Slide 12: PIGE in SIMNRA 7
Identical code basis to RBS/ERDA/NRA
Accuracy with default parameters ~10-5, can be improved by using fixed small step width
Straggling described by asymmetric straggling distributions (two-piece normal distribution) including 3rd moment
3rd moment for electronic energy loss straggling from free Coulomb scattering
3rd moment for multiple scattering from path length differences
Multiple scattering angle averaged mean path length in layers
Number of Gammas from Cross-section*Straggling
Multiple energies by scripting using OLE objects